BED RAIL POST INSTALLATION CHECK

| Service user’s name: |  | Assessor’s name: |  |
| --- | --- | --- | --- |
| Service user’s address: |  | Assessor’s contact number or email address: |  |

Type of bed rails loaned: Contract Trombone Contract High Sided Contract Sidhil Safe Sides Other Please specify:

| Are bumpers fitted? | Yes No |
| --- | --- |
| Are the bed rails high enough to take into account any increased mattress thickness or additional overlay? | Yes No |
| The ICES contract bed rails conform to British Standards and should have only been fitted to compatible beds – have they been set up correctly, ensuring any gaps are within recommended settings? (Please use the chart at the end of this form to check and record actual fitted measurements.) | Yes No |
| Does the gap between bumper/bed rails and edge of mattress present a risk to the service user? If the answer is yes, please take remedial action to resolve. | Yes No |
| Do the bed rails meet the identified needs? If no, please identify further actions required: | Yes No |
| Have the service user and carers been advised about the risks that bed rails may pose? | Yes No |
| Have you informed the service user and carers to report any adverse incidents and request urgent reassessment? | Yes No |
| Has a copy of the risk assessment been uploaded into Framework? | Yes No |
| Have ICES issued manufacturers equipment information/instructions? | Yes No |

| I have no concerns about the safety of the service user in using the bed rails |  |
| --- | --- |
| I have concerns about the safety of the service user using the bed rails and I have removed them Please give details of any other action taken |  |
| Date of installation check Signature of assessor |  |

GUIDANCE CHART

| Code | British Standard 2001 | Actual fitted measurements: |  |
| --- | --- | --- | --- |
| A | Should be a maximum of 120mm |  |  |
| B | Mattress depth including pressure mattress, if fitted, without compression |  |  |
| C | Should be a minimum of 220mm (from uncompressed mattress to top of rail) |  |  |
| D | Gap between headboard and end of bed rail no greater than 60mm Gap between footboard and end of bedrail should be less than 60mm or greater than 318mm | Head end: |  |
|  |  | Foot end: |  |
| E | Should be at least 2/3 of the bed |  |  |
| F | No requirements for bed length |  |  |

E

D

D

Bed head

No more than 60mm

Minimum 220 mm

A

C

F

B